- Tujuan/Objective

Untuk memastikan tata cara pemeriksaan bahan baku Hidroklorotiazid dilakukan dengan baik dan benar sesuai standar yang berlaku.

To guarantee that the method for testing Hydrochlorothiazide raw materials is followed appropriately and in accordance with applicable requirements.

- RuangLingkup/Scope

Laboratorium PT Kimia Farma Tbk. Plant Banjaran

PT Kimia Farma Tbk. Plant Banjaran laboratory.

- Penanggung Jawab/Responsibility

Analis Pengawasan Mutu bertanggung jawab melaksanakan pemeriksaan bahan baku Hidroklorotiazid sesuai dengan instruksi kerja ini.

The Quality Control Analyst is in charge of inspecting Hydrochlorothiazide raw materials in line with the job instructions.

- Referensi/Reference

Farmakope Indonesia VI, Tahun 2020

PT. Kimia Farma

- Alat dan Bahan/Equipment and Materials

  - Alat/Equipment

| 1. | Tabung Reaksi Test tube | 9. | Botol Timbang Weighing bottle |
| --- | --- | --- | --- |
| 2. | Gelas Kimia Beaker glass | 10. | Kurs Porselen Crucible |
| 3. | Tanur Furnace | 11. | Pipet seukuran Volume pipette |
| 4. | Oven Oven | 12. | Desikator Desiccator |
| 5. | Spektrofotometer IR IR Spectrophotometer | 13. | KCKT HPLC |
| 6. | Spektrofotometer UV-VIS UV-Vis Spectrophotometer | 14. | Timbangan analitik Analytical balance |
| 7. | Ultrasonik Ultrasonic bath | 15. | Tabung nessler Nessler tube |
| 8. | Saringan membran 0,20 µm Membrane filter 0,20 µm | 16. | Saringan syringe 0,20 µm Syringe filter 0,20 µm |

  - Bahan/Materials

| 1. | Natrium Hidroksida Sodium hydroxide | 9. | Asam klorida Hydrochloric acid |
| --- | --- | --- | --- |
| 2. | Baku kerja Hidrochlorothiazide Hydrochlorothiazide working standard | 10. | Asetonitril Acetonitrile |
| 3. | Baku kerja Klorotiazid Chlorothiazide working standard | 11. | Metanol Methanol |
| 4. | Senyawa sejenis A Benzotiadiazin Similar compound A Benzothiadiazine | 12. | Kalium Bromida P Potassium bromide |
| 5. | Air murni Purifying water | 13. | Asam Sulfat Sulfuric acid |
| 6. | Natrium fosfat monobasa Monobasic sodium phosphate | 14. | Perak nitrat LP Silver nitrate TS |
| 7. | Natrium fosfat Sodium phosphate | 15. | Asam nitrat 2 N; 2,5% 2 N; 2,5% nitric acid |
| 8. | Asam format Formic acid | 16. | Asam fosfat Phosphoric acid |

- Parameter Pemeriksaan/Inspection parameters

  - Pemerian/Description

    - Syarat/Requirement :

Sesuai dengan standar spesifikasi Hidroklorotiazid (S-QC-01-033/00)

Hydrochlorothiazide meets the specified specifications (S-QC-01-033/00)

    - Cara Kerja/ Procedure

Amati secara visual (Organoleptik)

Visually observe (organoleptic)

  - Kelarutan/Solubility

    - Syarat/Requirement :

Sesuai dengan standar spesifikasi Hidroklorotiazid (S-QC-01-033/00)

Hydrochlorothiazide meets the specified specifications (S-QC-01-033/00)

    - Cara Kerja/Procedure :

- Kelarutan dalam Metanol/Solubility in Methanol

- Larutkan 1 g zat uji dengan 100 ml Metanol dalam gelas kimia 100 ml.

Dissolve 1 g of the sample with 100 mL of Methanol in a 100 mL beaker.

- Amati pemerian larutan.

Observe description of solution.

- Kelarutan dalam Air murni/Solubility in water

- Larutkan 100 mg zat uji dengan 100 ml Air murni dalam gelas kimia 100 ml.

Dissolve 100 mg of the sample with 100 mL of Pure water in a 100 mL beaker.

- Amati pemerian larutan.

Observe description of solution.

- Kelarutan dalam HCl/Solubility in HCl

- Larutkan 1 g zat uji dengan 100 ml HCl dalam gelas kimia 100 ml.

Dissolve 1 g of the sample with 100 mL of HCl in 100 mL beaker.

- Amati pemerian larutan.

Observe description of solution.

  - Identifikasi/Identification

    - Syarat/Requirement :

Sesuai dengan standar spesifikasi Hidroklorotiazid (S-QC-01-033/00)

Hydrochlorothiazide meets the specified specifications (S-QC-01-033/00)

    - Cara Kerja/Procedure :

- Serapan Inframerah/Infrared absorption

- Zat Baku/Standard substance

- Campur dan gerus 8 mg baku kerja Hidroklorotiazid dengan 400 mg Kalium bromida kering.

Combine 8 mg of Hydrochlorothiazide working standard with 400 mg of dry potassium bromide in a mortar and pestle.

- Buat pellet Kalium bromida.

Make Potassium bromide pellets.

- Zat Uji/Sample substance

- Campur dan gerus 8 mg zat uji Hidroklorotiazid dengan 400 mg Kalium bromida kering.

Combine 8 mg of Hydrochlorothiazide sample with 400 mg of dry potassium bromide in a mortar and pestle.

- Buat pellet Kalium bromida.

Make Potassium bromide pellets.

- Pengukuran/Measurement

- Lakukan “scanning” menggunakan spektrofotometer IR pada panjang gelombang 400-4000 nm.

Use an IR spectrophotometer to do "scanning" at a wavelength of 400-4000 nm.

- Bandingkan spektra zat uji dengan zat baku.

Compare the spectra of the test substance with the standard substance.

- Amati kedua spektra tersebut, spektrum serapan inframerah zat uji menunjukkan maksimum hanya pada panjang gelombang yang sama seperti pada zat baku.

Observe the two spectra, the infrared absorption spectrum of the test substance shows a maximum only at the same wavelength as the standard substance.

- Serapan Ultraviolet/Ultraviolet absorption

- Larutan Baku/Standard solution

  - Timbang 10 mg baku kerja Hydrochlorothiazide, masukkan ke dalam labu ukur 50 ml.

Weigh 10 mg of working standard Hydrochlorothiazide, put it in a 50 mL volumetric flask.

  - Tambahkan beberapa ml Metanol P, sonikasi hingga larut.

Add a few mL of Methanol, sonicate until dissolved.

  - Encerkan dengan Metanol P hingga tanda batas, kocok hingga homogen.

Dilute with Methanol to mark the limit, shake until homogeneous.

  - Pipet 1 ml larutan, masukkan ke dalam labu ukur 20 ml.

Pipette 1 mL of the solution into a 20 mL volumetric flask.

  - Encerkan dengan Metanol P hingga tanda batas, kocok hingga homogen.

Dilute with Methanol to mark the limit, shake until homogeneous.

(Konsentrasi : 0,01 mg/ml)/(Concentration : 0,01 mg/mL)

- Larutan Uji/Sample solution

- Timbang 10 mg baku kerja Hydrochlorothiazide, masukkan ke dalam labu ukur 50 ml.

Weigh 10 mg of working standard Hydrochlorothiazide, put it in a 50 mL volumetric flask.

- Tambahkan beberapa ml Metanol P, sonikasi hingga larut.

Add a few mL of Methanol, sonicate until dissolved.

- Encerkan dengan Metanol P hingga tanda batas, kocok hingga homogen.

Dilute with Methanol to mark the limit, shake until homogeneous.

- Pipet 1 ml larutan, masukkan ke dalam labu ukur 20 ml.

Pipette 1 mL of the solution into a 20 mL volumetric flask.

- Encerkan dengan Metanol P hingga tanda batas, kocok hingga homogen.

Dilute with Methanol to mark the limit, shake until homogeneous.

(Konsentrasi : 0,01 mg/ml)/(Concentration : 0,01 mg/mL)

- Pengukuran/Measurement

- Ukur serapan larutan uji menggunakan Spektrofotometer UV-VIS pada Panjang gelombang 200 – 400 nm dengan Metanol sebagai blanko.

Measure the absorption of the test solution using a UV-VIS spectrophotometer at a wavelength of 200 - 400 nm with Methanol as a blank.

- Spektrum serapan ultraviolet menunjukkan serapan maksimum dan minimum pada Panjang gelombang yang sama dengan Hydochlorothiazide

The ultraviolet absorption spectrum shows the maximum and minimum absorption at the same wavelength as Hydrochlorothiazide.

  - Susut Pengeringan/Loss on Drying

    - Syarat/Requirement :

Sesuai dengan standar spesifikasi Hidroklorotiazid (S-QC-01-033/00)

Hydrochlorothiazide meets the specified specifications (S-QC-01-033/00)

    - Cara Kerja/Procedure :

- Tara botol timbang yang telah dikeringkan dalam oven selama 30 menit pada temperatur 105oC ± 2 oC (Wo).

Weigh the weighing bottle after it has been dried in the oven for 30 minutes at 105oC ± 2 oC (Wo).

- Bila zat uji berupa kristal besar, gerus secara cepat hingga ukuran partikel ± 2 mm.

If the test substance is a large crystal, destroy it rapidly up to a particle size of  ± 2 mm.

- Timbang seksama 1 - 2 g zat uji, masukkan ke dalam botol timbang yang telah ditara (Wa).

Weigh carefully 1-2 g of the test substance, put it in a weighing bottle that has been weighed (Wa).

- Masukkan ke dalam oven dengan suhu 105oC.

Put it in the oven at 105oC.

- Keringkan selama 1 jam.

Dry for 1 hours.

- Dinginkan dalam desikator selama 5-15 menit, kemudian timbang (Wb).

Cool in a desiccator for 5-15 minutes, then weigh (Wb).

    - Perhitungan/Calculation :

SP=Wa-WbWa-Wo x 100%

Keterangan/Note :

SP	= Susut pengeringan (%)

Loss on Drying (%)

Wa	= Berat botol timbang + zat uji sebelum pengeringan (gram)

Weight of weighing bottle + the substance before drying (grams)

Wb	= Berat botol timbang + zat uji setelah pengeringan (gram)

Weight of weighing bottle + the substance after drying (grams)

Wo	= Berat botol timbang kosong (gram)

Weight of empty weighing bottle (grams)

  - Sisa Pemijaran/Residue of Ignition

    - Syarat/Requirement :

Sesuai dengan standar spesifikasi Hidroklorotiazid (S-QC-01-033/00)

Hydrochlorothiazide meets the specified specifications (S-QC-01-033/00)

    - Cara Kerja/Procedure :

- Tara krus porselen yang telah dipijarkan selama 30 menit pada suhu 600oC ± 50oC (Wo).

Weigh the porcelain crucible that has been incandescent for 30 minutes at a temperature of 600oC ± 50oC (Wo).

- Timbang saksama 1-2 g zat uji ke dalam krus porselen (Wa).

Weigh 1-2 g of the test substance thoroughly into the porcelain crucible (Wa).

- Tambah 1 ml Asam sulfat P, panaskan di atas incinerator sampai zat uji mengarang sempurna, dinginkan.

Add 1 mL of Concentrated Sulfuric acid, heat over the incinerator until it is perfectly coaled, cool.

- Basahkan residu dengan 1 mL Asam sulfat P.

Wet the residue with 1 mL of Concentrated Sulfuric acid.

- Panaskan lagi hingga uap putih hilang.

Reheat until the white steam has dissipated.

- Pijarkan dalam tanur pada suhu 600oC ± 50oC hingga arang habis terbakar.

Ignite in furnace at  600oC ± 50oC until the charcoal is burned out.

- Dinginkan cawan dalam desikator, kemudian timbang (Wb).

Cool the crucible  in desiccator, and weigh (Wb).

- Jika jumlah sisa yang diperoleh lebih dari batas yang ditetapan, basahkan lagi sisa dengan 1 mL Asam sulfat P.

If the remaining amount is more than the limit set, rinse the remaining with 1 mL of Concentrated Sulfuric acid.

- Pijarkan dalam tanur pada suhu 600oC selama 30 menit.

Ignite in furnace at 600oC for 30 minutes.

- Dinginkan dalam desikator, kemudian timbang (Wb).

Cool in desiccator, then weigh (Wb).

    - Perhitungan/Calculation :

SP=Wb-WoWa-Wo×100%

Keterangan/Note :

SP		= Sisa pemijaran (%)

Residue on ignition (%)

Wo		= Berat krus porselen kosong (gram)

Weight of empty crucible (grams)

Wa		= Berat krus porselen + zat uji sebelum pemijaran (gram)

Weight of porcelain crucible + test substance before ignite (grams)

Wb		= Berat krus porselen + zat uji setelah pemijaran (gram)

Weight of porcelain crucible + test substance after ignite (grams)

  - Klorida/Chloride

    - Syarat/Requirement :

Sesuai dengan standar spesifikasi Hidroklorotiazid (S-QC-01-033/00)

Hydrochlorothiazide meets the specified specifications (S-QC-01-033/00)

    - Cara Kerja/Procedure :

- Larutan baku/Standard solutions

- Pipet 0,25 ml larutan HCl 0,020 N, masukkan ke dalam tabung Nessler.

Pipette 0.25 mL 0.020 N HCl solution, put it into the Nessler tube.

- Tambahkan 25 ml Air murni, kocok hingga homogen.

Add 25 mL of purifying water, shake until homogeneous.

- Larutan Uji/Sample solution

- Timbang 1,0 g zat uji, masukkan ke dalam tabung Nessler.

Weigh 1.0 g of the sample substance, put it into the Nessler tube.

- Tambahkan 80 mL Air murni, kocok hingga homogen.

Add 80 mL of purifying water, shake until homogeneous.

- Pengamatan/Observation

- Ke dalam masing – masing tabung Nessler/Into each Nessler tube :

- Tambahkan 1 mL HNO3 2N, kocok.

Add 1 mL 2 N HNO3, shake.

- Tambahkan 1 mL AgNO3 LP, kocok.

Add 1 mL AgNO3 TS, shake.

- Bandingkan kedua larutan tersebut.

Compare the two solutions.

- Amati kekeruhan larutan, intensitas kekeruhan larutan uji tidak lebih dari larutan baku.

Observe the turbidity of the solution, the intensity of the turbidity of the test solution is not more than the standard solution.

  - Logam Berat/Heavy Metal

    - Syarat/Requirement :

Sesuai dengan standar spesifikasi Hidroklorotiazid (S-QC-01-033/00)

Hydrochlorothiazide meets the specified specifications (S-QC-01-033/00)

    - Cara Kerja/Procedure :

- Pemeriksaan Timbal (Pb), Kadmium (Cd), dan Arsenik (As)

Examination of Lead (Pb), Cadmium (Cd), and Arsenic (As)

- Larutan Asam Nitrat 2,5%/Nitric Acid 2,5% solution

  - Ambil 500 ml air murni, masukkan ke dalam gelas kimia 1000 ml.

Fill a 1000 ml beaker halfway with 500 mL of fresh water.

  - Tambahkan 38,5 ml Asam Nitrat P secara perlahan, aduk hingga homogen.

Slowly pour in 38.5 mL of Nitric Acid  and mix until smooth.

  - Tambahkan air murni hingga tanda batas, aduk hingga homogen.

To indicate the limit, add pure water and stir until the mixture is uniform.

- Larutan Baku Multi-element Induk/Multi-element standard solution

- Pipet 2 ml ICP Multi-element Standard Solution IV 100 ppm, masukkan ke dalam labu ukur 20 ml.

Pipette 2 mL ICP Multi-element Standard Solution IV, 100 ppm, placed in a 20 mL volumetric flask.

- Tambahkan larutan Asam Nitrat 2,5% hingga tanda batas, kocok hingga homogen.

To indicate the limit, add 2.5% Nitric Acid solution and mix until homogenous.

- Saring menggunakan syringe filter 0,2 µm.

Filter using a 0.2 m syringe filter.

(Konsentrasi Pb, Cd, As: 10 ppm)/(Concentration of Pb, Cd, As: 10 ppm)

- Larutan Baku Multi-element 0,1 ppm/Multi-element standard solution 0,1 ppm

  - Pipet 0,5 ml Larutan Baku Multi-element Induk, masukkan ke dalam labu ukur   50 ml.

Pipette 0.5 mL of Multi-element Standard Solution into a volumetric flask with a 50 mL capacity.

  - Tambahkan larutan Asam Nitrat 2,5% hingga tanda batas, kocok hingga homogen.

To indicate the limit, add 2.5% Nitric Acid solution and mix until homogenous.

  - Saring menggunakan syringe filter 0,2 µm.

Filter using a 0.2 m syringe filter.

(Konsentrasi Pb, Cd, As: 0,1 ppm)/(Concentration of Pb, Cd, As: 0,1 ppm)

- Larutan Baku Multi-element 0,3 ppm/Multi-element standard solution 0,3 ppm

- Pipet 1,5 ml Larutan Baku Multi-element Induk, masukkan ke dalam labu ukur   50 ml.

Pipette 1.5 mL of Multi-element Standard Solution into a volumetric flask with a 50 mL capacity.

- Tambahkan larutan Asam Nitrat 2,5% hingga tanda batas, kocok hingga homogen.

To indicate the limit, add 2.5% Nitric Acid solution and mix until homogenous.

- Saring menggunakan syringe filter 0,2 µm.

Filter using a 0.2 m syringe filter.

(Konsentrasi Pb, Cd, As: 0,3 ppm)/(Concentration of Pb, Cd, As: 0,5 ppm)

- Larutan Baku Multi-element 0,5 ppm/Multi-element standard solution 0,5 ppm

- Pipet 2,5 ml Larutan Baku Multi-element Induk, masukkan ke dalam labu ukur   50 ml.

Pipette 2.5 mL of Multi-element Standard Solution into a volumetric flask with a 50 mL capacity.

- Tambahkan larutan Asam Nitrat 2,5% hingga tanda batas, kocok hingga homogen.

To indicate the limit, add 2.5% Nitric Acid solution and mix until homogenous.

- Saring menggunakan syringe filter 0,2 µm.

Filter using a 0.2 m syringe filter.

(Konsentrasi Pb, Cd, As: 0,5 ppm)/(Concentration of Pb, Cd, As: 0,5 ppm)

- Larutan Uji/Sample solution

- Timbang 2 g zat uji, masukkan ke dalam labu ukur 20 ml.

Weigh 2 g of the test substance, put it in a 20 mL volumetric flask.

- Tambahkan larutan Asam Nitrat 2,5%, sonikasi hingga larut.

Add 2.5% nitric acid solution, sonicate until dissolved.

- Tambahkan larutan Asam Nitrat 2,5% hingga tanda batas volume, kocok hingga homogen.

To indicate the limit, add 2.5% Nitric Acid solution and mix until homogenous.

- Endapkan hingga endapan dan filtrat terpisah.

Precipitate until the filtrate and precipitate are separated.

- Saring filtrat menggunakan syringe filter 0,2 µm.

Filter the filtrate using a 0.2 m syringe filter.

- Pemeriksaan Raksa/Examination of Mercury

- Larutan Asam Nitrat 13%/Nitric Acid 13% solution

  - Ambil 500 ml air murni, masukkan ke dalam gelas kimia 1000 ml.

Fill a 1000 mL beaker halfway with 500 mL of fresh water.

  - Tambahkan 200 ml Asam Nitrat pekat secara perlahan, aduk hingga homogen.

Slowly pour in 200 mL of Nitric Acid and mix until smooth.

  - Tambahkan air murni hingga tanda batas, aduk hingga homogen.

To indicate the limit, add pure water and stir until the mixture is uniform.

- Larutan Baku Raksa Induk/Mercury standard solution

  - Pipet 10 ml Hg Standard For ICP 1000 ppm, masukkan ke dalam labu ukur 100 ml.

Pipette 10 mL of Hg Standard For ICP 1000 ppm into a volumetric flask with a 100 mL capacity.

  - Tambahkan larutan Asam Nitrat 2,5% hingga tanda batas, kocok hingga homogen.

To indicate the limit, add 2.5% Nitric Acid solution and mix until homogenous.

  - Saring menggunakan syringe filter 0,2 µm.

Filter using a 0.2 m syringe filter.

(Konsentrasi Hg: 100 ppm)/(Concentration of Hg: 100 ppm)

- Larutan Baku Raksa 1 ppm/Mercury standard solution 1 ppm

  - Pipet 0,5 ml Larutan Baku Raksa Induk, masukkan ke dalam labu ukur 50 ml.

Pipette 0.5 mL of Mercury Standard Solution into a volumetric flask with a 50 mL capacity.

  - Tambahkan larutan Asam Nitrat 13% hingga tanda batas, kocok hingga homogen.

To indicate the limit, add 13% Nitric Acid solution and mix until homogenous.

  - Saring menggunakan syringe filter 0,2 µm.

Filter using a 0.2 m syringe filter.

(Konsentrasi Hg: 1 ppm)/(Concentration of Hg: 1 ppm)

- Larutan Baku Raksa 2 ppm/Mercury standard solution 2 ppm

  - Pipet 1 ml Larutan Baku Raksa Induk, masukkan ke dalam labu ukur   50 ml.

Pipette 1 mL of Mercury Standard Solution into a volumetric flask with a 50 mL capacity.

  - Tambahkan larutan Asam Nitrat 13% hingga tanda batas, kocok hingga homogen.

To indicate the limit, add 13% Nitric Acid solution and mix until homogenous.

  - Saring menggunakan syringe filter 0,2 µm.

Filter using a 0.2 m syringe filter.

(Konsentrasi Hg: 3 ppm)/(Concentration of Hg: 3 ppm)

- Larutan Baku Raksa 3 ppm/Mercury standard solution 3 ppm

  - Pipet 1,5 ml Larutan Baku Raksa Induk, masukkan ke dalam labu ukur   50 ml.

Pipette 1.5 mL of Mercury Standard Solution into a volumetric flask with a 50 mL capacity.

  - Tambahkan larutan Asam Nitrat 13% hingga tanda batas, kocok hingga homogen.

To indicate the limit, add 13% Nitric Acid solution and mix until homogenous.

  - Saring menggunakan syringe filter 0,2 µm.

Filter using a 0.2 m syringe filter.

(Konsentrasi Hg: 2 ppm)/(Concentration of Hg: 2 ppm)

- Larutan Baku Raksa 4 ppm/Mercury standard solution 4 ppm

  - Pipet 2 ml Larutan Baku Multi-element Induk, masukkan ke dalam labu ukur   50 ml.

Pipette 2 mL of Mercury Standard Solution into a volumetric flask with a 50 mL capacity.

  - Tambahkan larutan Asam Nitrat 13% hingga tanda batas, kocok hingga homogen.

To indicate the limit, add 13% Nitric Acid solution and mix until homogenous.

  - Saring menggunakan syringe filter 0,2 µm.

Filter using a 0.2 m syringe filter.

(Konsentrasi Hg: 4 ppm)/(Concentration of Hg: 4 ppm)

- Larutan Baku Raksa 5 ppm/Mercury standard solution 5 ppm

  - Pipet 1 ml Larutan Baku Multi-element Induk, masukkan ke dalam labu ukur   50 ml.

Pipette 1 mL of Mercury Standard Solution into a volumetric flask with a 50 mL capacity.

  - Tambahkan larutan Asam Nitrat 13% hingga tanda batas, kocok hingga homogen.

To indicate the limit, add 13% Nitric Acid solution and mix until homogenous.

  - Saring menggunakan syringe filter 0,2 µm.

Filter using a 0.2 m syringe filter.

(Konsentrasi Hg: 5 ppm)/(Concentration of Hg: 5 ppm)

- Larutan Uji/Sample solution

  - Timbang 2 g zat uji, masukkan ke dalam labu ukur 20 ml.

Weigh 2 g of the test substance, put it in a 20 mL volumetric flask.

  - Larutkan dengan Asam Nitrat 13%, sonikasi hingga larut.

Solve with 2.5% nitric acid solution, sonicate until dissolved.

  - Tambahkan Asam Nitrat 13% hingga tanda batas, kocok hingga homogen.

To indicate the limit, add 13% Nitric Acid solution and mix until homogenous.

  - Endapkan hingga endapan dan filtrat terpisah.

Precipitate until the filtrate and precipitate are separated.

  - Saring filtrat menggunakan syringe filter 0,2 µm.

Filter the filtrate using a 0.2 m syringe filter.

      - Pengukuran/Measurement

Lakukan pemeriksaan sesuai dengan Instruksi Kerja Pengoperasian Alat Inductively Coupled Plasma.

Follow the Inductively Coupled Plasma Equipment Work Instructions for inspecting.

  - Senyawa Sejenis/Similar Compound

    - Syarat/Requirement :

Sesuai dengan standar spesifikasi Hidroklorotiazid (S-QC-01-033/00)

Hydrochlorothiazide meets the specified specifications (S-QC-01-033/00)

    - Cara Kerja/Procedure :

- Larutan Natrium Fosfat/Sodium phosphate solution

- Timbang 2,76 g Natrium fosfat monobasa (NaH2PO4.H2O), masukkan ke dalam labu ukur 1000 ml.

Weigh 2.76 g of monobasic sodium phosphate (NaH2PO4.H2O), put it in a 1000 mL volumetric flask.

- Tambahkan 990 ml Air murni, aduk hingga homogen.

Add 990 mL of pure water, stir until homogeneous.

- Atur pH hingga 2,7 ± 0,1 dengan penambahan Asam fosfat P.

With the addition of Phosphoric Acid P, adjust the pH to 2.7 ± 0.1.

- Tambahkan Air murni hingga batas volume, kocok hingga homogen.

Add pure water to the volume limit, shake until homogeneous.

- Larutan Pengencer/Diluent

- Buat campuran larutan Natrium fosfat : Asetonitril P (7:3), masukkan ke dalam gelas kimia 1000 ml.

Make a mixture of Sodium phosphate solution : Acetonitrile P (7:3), put it in a 1000 mL beaker.

- Aduk hingga homogen.

Shake until homogeneous.

- Fase Gerak (Gunakan variasi campuran larutan A dan larutan B)

Mobile phase (Use a variant of a mixture of solutions A and B)

- Larutan A/Solution A

- Buat campuran Asetonitril P : Metanol P (3:1), masukkan ke dalam reagen fase gerak.

Make a mixture of Acetonitrile P: Methanol P (3:1), put it in the mobile phase reagent.

- Aduk hingga homogen.

Shake until homogeneous.

- Saring larutan dengan membran filter 0,20 µm, awaudarakan selama 10 menit.

Filter the solution through a 0.20 µm filter membrane, and allow to air for 10 minutes.

- Larutan B/Solution B

- Buat campuran Asam format anhidrat dalam Air murni (5 dalam 1000), masukkan ke dalam reagen fase gerak.

Make a mixture of anhydrous formic acid in purified water (5 in 1000), put it in the mobile phase reagent.

- Aduk hingga homogen.

Shake until homogeneous.

- Saring larutan dengan membran filter 0,20 µm, awaudarakan selama 10 menit.

Filter the solution through a 0.20 µm filter membrane, and allow to air for 10 minutes.

| Waktu (menit) Time (minute) | Larutan A (%) Solution A (%) | Larutan B (%) Solution B (%) |
| --- | --- | --- |
| 0 | 3 | 97 |
| 0-5 | 3 | 97 |
| 5-14 | 3 → 36 | 97 → 64 |
| 14-18 | 36 → 3 | 64 → 97 |
| 18-20 | 3 | 97 |

d. Sistem Kromatografi/Chromatography System

Kolom		: Inertsil ODS-3; C-18 L1; 4,6 x 150 mm; 5 µm

Column		: Inertsil ODS-3; C-18 L1; 4,6 x 150 mm; 5 µm

Detektor	: UV 275 nm

Detector		: UV 275 nm

Flow rate	: 1,5 ml/menit

Flow rate		: 1,5 mL/minute

Volume Injeksi	: 10 μl

Injection volume	: 10 μL

Suhu kolom	: 35°C

Column temperature	: 35°C

e. Kriteria Penerimaan/Acceptance Criteria

| No | Parameter/Parameter | Syarat/Requirement |
| --- | --- | --- |
| 1 | Luas Area Area | Nilai Simpangan Baku Relatif (SBR) : Relative standard deviation ≤ 2,0 % - Hydrochlorothiazide : ≤ 2,0 % - Senyawa sejenis A Benzotiadiazine : ≤ 5,0 % - Klorotiazid : ≤ 5,0 % |
| 2 | Waktu Retensi (tr) Retention time (tr) | Nilai Simpangan Baku Relatif (SBR) ≤ 2,0 % Relative standard deviation ≤ 2,0 - Hydrochlorothiazide : ≤ 2,0 % - Senyawa sejenis A Benzotiadiazine : ≤ 5,0 % - Klorotiazid : ≤ 5,0 % |
| 3 | Tailing Faktor (tf) Tailing factor (tf) | Tf ≤ 2,0 |
| 4 | Lempeng Teoritis (N) Theoretical plate (N) | N ≥ 2000 |
| 5 | Resolusi (R) Resolution (R) | - Antara puncak senyawa sejenis A Benzotiadiazin dan puncak Klorotiazid ≥ 2,0 Between the peaks of benzothiadiazine A-like compounds and Chlorothiazide ≥ 2.0 - Antara puncak Klorotiazid dan puncak Hydrochlorotiazide ≥ 1,5 Between Chlorothiazide peak and Hydrochlorothiazide peak ≥ 1.5 |

f. Larutan Batas Kuantitasi/Quantitation limit solution

- Timbang 8 mg baku kerja Hydrochlorotiazid, masukkan ke dalam labu ukur 100 ml.

Weigh 8 mg of the working standard Hydrochlorothiazide, put it in a 100 ml volumetric flask.

- Tambahkan beberapa ml larutan pengencer, sonikasi selama 10 menit.

Add a few mL of diluent, sonicate for 10 minutes.

- Encerkan dengan larutan pengencer hingga tanda batas, homogenkan (larutan A).

Dilute with diluent to the mark, homogenize (solution A).

- Pipet 1 ml larutan A, masukkan ke dalam labu ukur 50 ml.

Pipette 1 ml of solution A into a 50 mL volumetric flask.

- Encerkan dengan larutan pengencer hingga tanda batas, homogenkan (larutan B).

Dilute with diluent to the mark, homogenize (solution B).

- Pipet 1 ml larutan B, masukkan ke dalam labu ukur 10 ml.

Pipette 1 ml of solution B into a 10 mL volumetric flask.

- Encerkan dengan larutan pengencer hingga tanda batas, homogenkan.

Dilute with diluent to the mark, homogenize.

- Saring dengan syringe filter 0,20 μm, tampung filtrat ke dalam vial.

Filter using a 0.20 µm syringe filter and transfer the filtrate to a vial.

(Konsentrasi Hydrochlorotiazid : 0,00016 mg/ml = 0,16 μg/ml)

(Concentration Hydrochlorothiazide : 0,00016 mg/mL = 0,16 μg/mL)

g. Larutan Uji/Sample Solution

- Timbang 32 mg zat uji, masukkan ke dalam labu ukur 100 ml.

Weigh 32 mg of the test substance, put it in a 100 mL volumetric flask.

- Tambahkan beberapa ml larutan pengencer, sonikasi selama 10 menit.

Add a few mL of diluent, sonicate for 10 minutes.

- Encerkan dengan larutan pengencer hingga tanda batas, homogenkan.

Dilute with diluent to the mark, homogenize.

- Saring dengan syringe filter 0,20 μm, tampung filtrat ke dalam vial.

Filter using a 0.20 μm syringe filter and transfer the filtrate to a vial.

(Konsentrasi : 0,32 mg/ml)/(Concentration : 0,32 mg/mL)

h. Perhitungan/Calculation :

C= LicLsc x 100%

Keterangan/Note :

| C | = | Kadar cemaran yang terdapat dalam bahan baku Hydrochlorothiazide (%) Contaminants contained in the raw material Hydrochlorothiazide (%) |
| --- | --- | --- |
| Lic | = | Perbandingan luas area masing-masing cemaran terhadap faktor respon Comparison of each contamination's region to the response components |
| Lsc | = | Jumlah perbandingan semua luas area terhadap semua faktor respon The total of all comparisons between all geographical areas and all response components |

  - Kadar/Assay

    - Syarat/Requirement :

Sesuai dengan standar spesifikasi Hidroklorotiazid (S-QC-01-033/00)

Hydrochlorothiazide meets the specified specifications (S-QC-01-033/00)

    - Cara Kerja/Procedure :

- Larutan Natrium Fosfat/Sodium phosphate solution

- Timbang 2,76 g Natrium fosfat monobasa (NaH2PO4.H2O), masukkan ke dalam labu ukur 1000 ml.

Weigh 2.76 g of monobasic sodium phosphate (NaH2PO4.H2O), put it in a 1000 mL volumetric flask.

- Tambahkan 990 ml Air murni, aduk hingga homogen.

Add 990 mL of pure water, stir until homogeneous.

- Atur pH hingga 2,7 ± 0,1 dengan penambahan Asam fosfat P.

With the addition of Phosphoric Acid P, adjust the pH to 2.7 ± 0.1.

- Tambahkan Air murni hingga batas volume, kocok hingga homogen.

Add pure water to the volume limit, shake until homogeneous.

- Larutan Pengencer/Diluent

- Buat campuran larutan Natrium fosfat : Asetonitril P (7:3), masukkan ke dalam gelas kimia 1000 ml.

Make a mixture of Sodium phosphate solution : Acetonitrile P (7:3), put it in a 1000 mL beaker.

- Aduk hingga homogen.

Shake until homogeneous.

- Fase Gerak (Gunakan variasi campuran larutan A dan larutan B)

Mobile phase (Use a variant of a mixture of solutions A and B)

- Larutan A/Solution A

- Buat campuran Asetonitril P : Metanol P (3:1), masukkan ke dalam reagen fase gerak.

Make a mixture of Acetonitrile P: Methanol P (3:1), put it in the mobile phase reagent.

- Aduk hingga homogen.

Shake until homogeneous.

- Saring larutan dengan membran filter 0,20 µm, awaudarakan selama 10 menit.

Filter the solution through a 0.20 µm filter membrane, and allow to air for 10 minutes.

- Larutan B/Solution B

- Buat campuran Asam format anhidrat dalam Air murni (5 dalam 1000), masukkan ke dalam reagen fase gerak.

Make a mixture of anhydrous formic acid in purified water (5 in 1000), put it in the mobile phase reagent.

- Aduk hingga homogen.

Shake until homogeneous.

- Saring larutan dengan membran filter 0,20 µm, awaudarakan selama 10 menit.

Filter the solution through a 0.20 µm filter membrane, and allow to air for 10 minutes.

| Waktu (menit) Time (minute) | Larutan A (%) Solution A (%) | Larutan B (%) Solution B (%) |
| --- | --- | --- |
| 0 | 3 | 97 |
| 0-5 | 3 | 97 |
| 5-14 | 3 → 36 | 97 → 64 |
| 14-18 | 36 → 3 | 64 → 97 |
| 18-20 | 3 | 97 |

- Sistem Kromatografi/Chromatography System

Kolom		: Inertsil ODS-3; C-18 L1; 4,6 x 150 mm; 5 µm

Column		: Inertsil ODS-3; C-18 L1; 4,6 x 150 mm; 5 µm

Detektor	: UV 275 nm

Detector		: UV 275 nm

Flow rate	: 1,5 ml/menit

Flow rate		: 1,5 mL/minute

Volume Injeksi	: 10 μl

Injection volume	: 10 μL

Suhu kolom	: 35°C

Column temperature	: 35°C

- Kriteria Penerimaan/Acceptance Criteria

| No | Parameter/Parameter | Syarat/Requirement |
| --- | --- | --- |
| 1 | Luas Area Area | Nilai Simpangan Baku Relatif (SBR) ≤ 2,0 % Relative standard deviation ≤ 2,0 % |
| 2 | Waktu Retensi (tr) Retention time (tr) | Nilai Simpangan Baku Relatif (SBR) ≤ 2,0 % Relative standard deviation ≤ 2,0 |
| 3 | Tailing Faktor (tf) Tailing factor (tf) | Tf ≤ 2,0 |
| 4 | Lempeng Teoritis (N) Theoretical plate (N) | N ≥ 2000 |
| 5 | Resolusi (R) Resolution (R) | - Antara puncak senyawa sejenis A Benzotiadiazin dan puncak Klorotiazid ≥ 2,0 Between the peaks of benzothiadiazine A-like compounds and Chlorothiazide ≥ 2.0 - Antara puncak Klorotiazid dan puncak Hydrochlorotiazide ≥ 1,5 Between Chlorothiazide peak and Hydrochlorothiazide peak ≥ 1.5 |

- Larutan Uji Kesesuaian Sistem/Reference Standard Solution

- Larutan A/Solution A

- Timbang 16 mg baku kerja Hydrochlorothiazide, masukkan ke dalam labu ukur 50 ml.

Weigh 16 mg of working standard Hydrochlorothiazide, put it in a 50 mL volumetric flask.

- Tambahkan larutan pengencer beberapa ml, sonikasi selama 10 menit.

Add a few mL of diluent, sonicate for 10 minutes.

- Encerkan dengan larutan pengencer hingga tanda batas, homogenkan.

Dilute with diluent to the mark, homogenize.

- Pipet 1 ml larutan, masukkan ke dalam labu ukur 100 ml.

Pipette 1 mL of the solution into a 100 mL volumetric flask.

- Encerkan dengan larutan pengencer hingga tanda batas, homogenkan.

Dilute with diluent to the mark, homogenize.

(Konsentrasi : 3,2 µg/ml)/(Concentration : 3,2 µg/mL)

- Larutan B/Solution B

- Timbang 3,2 mg baku kerja Klorotiazid, masukkan ke dalam labu ukur 100 ml.

Weigh 3.2 mg of chlorothiazide working standard, put it in a 100 mL volumetric flask.

- Tambahkan beberapa ml larutan pengencer, sonikasi selama 10 menit.

Add a few mL of diluent, sonicate for 10 minutes.

- Encerkan dengan larutan pengencer hingga tanda batas, homogenkan.

Dilute with diluent to the mark, homogenize.

(Konsentrasi : 32 µg/ml)/(Concentration : 3,2 µg/mL)

- Larutan C/Solution C

- Timbang 3,2 mg Senyawa Sejenis A Benzotiadiazin, masukkan ke dalam labu ukur 100 ml.

Weigh 3.2 mg of Benzothiadiazine Similar Compound A, put it in a 100 mL volumetric flask.

- Tambahkan beberapa ml larutan pengencer, sonikasi selama 10 menit.

Add a few mL of diluent, sonicate for 10 minutes.

- Encerkan dengan larutan pengencer hingga tanda batas, homogenkan.

Dilute with diluent to the mark, homogenize.

(Konsentrasi : 32 µg/ml)/(Concentration : 3,2 µg/mL)

- Larutan Uji Kesesuaian Sistem/Reference standard solution

- Pipet larutan A sebanyak 1 ml, masukkan ke dalam labu ukur 10 ml.

Pipette 1 mL of solution A into a 10 mL volumetric flask.

- Pipet larutan B sebanyak 1 ml, masukkan ke dalam labu ukur 10 ml yang sama.

Pipette 1 mL of solution B into the same 10 mL volumetric flask.

- Pipet larutan C sebanyak 1 ml, masukkan ke dalam labu ukur 10 ml yang sama.

Pipette 1 mL of solution C into the same 10 mL volumetric flask.

- Encerkan dengan larutan pengencer hingga tanda batas, homogenkan.

Dilute with diluent to the mark, homogenize.

- Saring dengan syringe filter 0,20 μm, tampung filtrat ke dalam vial.

Filter using a 0.20 µm syringe filter and transfer the filtrate to a vial.

(Konsentrasi Hydrochlorothiazide : 0,32 µg/ml, Klorotiazid : 3,2 µg/ml, dan senyawa sejenis A Benzotiadiazin : 3,2 µg/ml)

(Concentration Hydrochlorothiazide : 0,32 µg/mL, Chlorothiazide : 3,2 µg/mL, and Benzothiadiazine Similar Compound A : 3,2 µg/mL)

- Larutan Baku/Standard solution

- Timbang 32,0 mg baku kerja Hydrochlorothiazide masukan ke dalam labu ukur 20 ml.

Weigh 32.0 mg Hydrochlorothiazide working standard put in 20 mL volumetric flask.

- Tambahkan beberapa ml larutan pengencer, sonikasi selama 10 menit.

Add a few mL of diluent, sonicate for 10 minutes.

- Encerkan dengan larutan pengencer hingga tanda batas, homogenkan.

Dilute with diluent to the mark, homogenize.

- Pipet 2 ml larutan ke dalam labu ukur 10 ml.

Pipette 2 mL of the solution into a 10 ml volumetric flask.

- Encerkan dengan larutan pengencer hingga tanda batas, homogenkan.

Dilute with diluent to the mark, homogenize.

- Saring dengan syringe filter 0,20 μm, tampung filtrat ke dalam vial.

Filter using a 0.20 µm syringe filter and transfer the filtrate to a vial.

(Konsentrasi : 0,32 mg/ml)/(Concentration : 0,32 mg/mL)

- Larutan Uji/Sample Solution

- Timbang 32 mg zat uji, masukkan ke dalam labu ukur 100 ml.

Weigh 32 mg of the test substance, put it in a 100 mL volumetric flask.

- Tambahkan beberapa ml larutan pengencer, sonikasi selama 10 menit.

Add a few mL of diluent, sonicate for 10 minutes.

- Encerkan dengan larutan pengencer hingga tanda batas, homogenkan.

Dilute with diluent to the mark, homogenize.

- Saring dengan syringe filter 0,20 μm, tampung filtrat ke dalam vial.

Filter using a 0.20 μm syringe filter and transfer the filtrate to a vial.

(Konsentrasi : 0,32 mg/ml)/(Concentration : 0,32 mg/mL)

- Perhitungan/Calculation :

C= LuLs x Cs Cu x 100%(100-SP)x B%

Keterangan/Note :

| C | = | Kadar Hydrochlorothiazide yang telah diperhitungkan terhadap baku kerja (%) Hydrochlorothiazide assay that have been calculated against work standards (%) |
| --- | --- | --- |
| Lu | = | Luas area larutan uji The area of the test solution |
| Ls | = | Luas area larutan baku The area of the standard solution |
| Cs | = | Konsentrasi larutan baku (mg/ml) Concentration of standard substance (mg/mL) |
| Cu | = | Konsentrasi larutan uji (mg/ml) Concentration of sample substance (mg/mL) |
| B% | = | Kadar baku kerja (%) Working standard assay (%) |
| SP | = | Susut pengeringan (%) Loss on drying (%) |

- Daftar Lampiran/ List of Attachment

| Lampiran ke- | No. Dokumen | Judul | Tanggal Berlaku |
| --- | --- | --- | --- |
| - | - | - | - |

- Sejarah RevisiProsedur/ Revision History of Procedure

| Kode Dokumen | Revisi ke- | Tanggal Berlaku Dokumen | Point Yang Berubah | # Alasan Perubahan |
| --- | --- | --- | --- | --- |
| I-QC-01-052 | 00 |  | - | Dokumen Baru |

- Review Dokumen/Document Review

| Review Ke | Tanggal | Hasil Review | Reviewer | Paraf |
| --- | --- | --- | --- | --- |
| - | - | - | - | - |

- Daftar Distribusi Dokumen/ Distribution List of Document

Asli : SQC

Controlled Copy:

AMQA

AMQC

SQTS

SCRP