- Objective

To guarantee that the method for testing Molnupiravir raw materials is followed appropriately and in accordance with applicable requirements.

- Scope

PT. Kimia Farma, Tbk. - Plant Banjaran Laboratory

- Responsibility

The Quality Control Analyst responsible for inspecting Molnupiravir raw materials in accordance with the job instructions.

- Definition

-

- Reference

In House Method (Adopted from API manufacturer Shanghai Desano)

Modification USP 46 NF 41 (2023)

- Tool and Material

  - Equipment :

| 1. | Burette | 17. | Magnetic Stirrer |
| --- | --- | --- | --- |
| 2. | Crucible | 18. | Membran Filter 0,2 µm |
| 3. | Desiccator | 19. | Syringe Filter 0,2 μm |
| 4. | Furnace | 20. | Spektrofotometer IR |
| 5. | Beaker Glass | 21. | Crucible Tong |
| 6. | Measuring Cylinder | 22. | Drop Pipette |
| 7. | Hot Plate | 23. | Analytical Balance |
| 8. | HPLC | 24. | Ultrasonic Bath |
| 9. | Karl Fischer | 25. | HPLC Vials |
| 10. | Polarimeter | 26. | Crimp Tool |
| 11. | Vacuum Pump | 27. | Gas Chromatography (GC) |
| 12. | Column Xterra RP18 25 cm x 4.6 mm; 5 m | 28. | GC Column: DB-1 60 m x 0.53 mm; 3 µm |
| 13. | Polarimeter Tube | 29. | GC Column: HP-5 30 m x 0.53 mm; 5 µm |
| 14. | Inductively Coupled Plasma (ICP) | 30. | Micropipette |
| 15. | Volumetric Flask | 31. | Micropipette Tips |
| 16. | Mortar and Pestle | 32. | Vial Headspace |

  - Materials :

| 1. | Purified Water | 13. | ICP Multi-element Standard Solution IV 100 ppm |
| --- | --- | --- | --- |
| 2. | Orthophosphoric Acid | 14. | Sodium Tartrate Dihydrate |
| 3. | Sulfuric Acid | 15. | Benzene |
| 4. | Acetonitrile | 16. | Dichloromethane |
| 5. | Ice Cube | 17. | Dimetylsulfoxide |
| 6. | Dried Potassium Bromide | 18. | Ethyl Acetate |
| 7. | Methanol | 19. | Isopropyl Alcohol |  |
| 8. | Nitric Acid | 20. | Isopropyl Acetate |  |
| 9. | Solvent Karl Fischer | 21. | N,N-dimethylformamide |  |
| 10. | Titrant Karl Fischer | 22. | Pyridine |  |
| 11. | Impurity Standard D | 23. | Triethylamine |  |
| 12. | 2-methyl Tetrahydrofuran |  |  |  |

- Work Instruction

  - Description

    - Requirement :

Molnupiravir Raw Material specification are met a requirement (S-QC-01-280).

    - Procedure :

Visually observe.

  - Identification

    - Requirement :

Molnupiravir Raw Material specification are met a requirement (S-QC-01-280).

    - Procedure :

- Infrared Absorption

- Standard Substance

- Mix and grind 8 mg of Molnupiravir working standard with 400 mg of dry Potassium Bromide in a mortar and pestle.

- Make Potassium Bromide pellets.

- Sample Substance

- Mix and grind 8 mg of Molnupiravir sample with 400 mg of dry Potassium Bromide in a mortar and pestle.

- Make Potassium Bromide pellets.

- Measurement

- Use an IR spectrophotometer to do "scanning" at a wavelength of 400 - 4000 nm.

- Compare the spectra of the test substance with the standard substance.

- Observe the two spectra, the infrared absorption spectrum of the test substance shows a maximum only at the same wavelength as the standard substance.

- HPLC Methode

The retention time at the peak of the Molnupiravir test substance was the same as the retention time at the peak of the Molnupiravir standard substance. The work of identification parameters follows the way the assay parameters inpection.

  - Water Content

    - Requirement :

Molnupiravir Raw Material specification are met a requirement (S-QC-01-280).

    - Procedure :

Follow the Karl Fischer Equipment Work Instructions when performing inspections.

  - Optical Rotation

    - Requirement :

Molnupiravir Raw Material specification are met a requirement (S-QC-01-280).

    - Procedure :

- Weigh 250 mg of sample substance, put it into 25 mL volumetric flask.

- Add about 15 mL of Methanol P, sonicate until dissolved.

- Dilute to volume with the same solvent, shake until homogeneous.

- Soak the volumetric flask with ice to a temperature of 25oC.

- Follow the Polarimeter Equipment Work Instructions when performing the inspection at temperature of 25oC.

    - Calculation :

α = P × 100 × Fp × 100%Wu × L × 100 × (100% -KA)

α = P × 100 × Fp × 100%Wu × L × 100 × (100% -KA)

Note :

| α | = | Specific rotation with respect to wavelength (deg) |
| --- | --- | --- |
| W | = | Weight of sample substance (g) |
| L | = | Polarimeter tube length (1 dm) |
| Fp | = | Dilution factor (mL) |
| P | = | Polarimeter result reading (deg) |
| KA | = | Water Content (%) |

  - Residue of Ignition

    - Requirement :

Molnupiravir Raw Material specification are met a requirement (S-QC-01-280).

    - Procedure :

- Incandescent crucible at a temperature of 600 ± 50oC in the Furnace, for 30 minutes.

- Cool the crucible in a desiccator before weighing it (Wo).

- Weigh accurately 1 - 2 g of sample substance, put into crucible (Wa).

- Add 1 mL of concentrated Sulfuric Acid.

- Heat until sample substance is carbonized perfectly, then allow to cool.

- Moisten the residue using 1 mL of concentrated Sulfuric Acid.

- Reheat until the white haze has gone.

- Incandescent in the Furnace at a temperature of 600 ± 50 oC until the charcoal burns out.

- Allow crucible to cool in desiccators, weigh carefully (Wb).

- If the amount of residue obtained exceeds the specified limit, moisten the residue again with 1 mL of Sulfuric Acid P.

- Reheat until the white haze has gone.

- Incandescent in the Furnace at a temperature of 600 ± 50 oC for 30 minutes.

- Allow crucible to cool in desiccators, weigh carefully (Wb).

    - Calculation :

SPn= Wb-WoWa-Wo × 100%

SPn= Wb-WoWa-Wo × 100%

Note :

SPn	= Residue of ignition (%)

Wo	= Weight of empty crucible (grams)

Wa	= Weight of crucible + sample substance before ignition (grams)

Wb	= Weight of crucible + sample substance after ignition (grams)

  - Heavy Metal

    - Requirement :

Molnupiravir Raw Material specification are met a requirement (S-QC-01-280).

    - Procedure :

- Solution Preparation

- 2.5% Nitric Acid Solution

      - Fill a 1000 mL beaker glass halfway with 500 mL of Purified Water.

- Add 38.5 mL of Nitric Acid P slowly; stir until homogeneous.

- Add 461.5 mL of Purified Water; stir until homogeneous.

- Multi-element Standard Solution

- Pipette 2 mL ICP Multi-element Standard Solution IV 100 ppm, put into a 20 mL volumetric flask.

- Add 2.5% Nitric Acid solution to volume and shake until homogenous.

- Filter using a 0.2 µm syringe filter.

(Concentration of Pb: 10 ppm)

- Multi-element Standard Solution 0.1 ppm

- Pipette 0.5 mL Multi-element Standard Solution; put into a 50 mL volumetric flask.

- Add 2.5% Nitric Acid solution to volume and shake until homogenous.

- Filter using a 0.2 µm syringe filter.

(Concentration of Pb: 0.1 ppm)

- Multi-element Standard Solution 0.3 ppm

- Pipette 1.5 mL Multi-element Standard Solution; put into a 50 mL volumetric flask.

- Add 2.5% Nitric Acid solution to volume; and shake until homogenous.

- Filter using a 0.2 µm syringe filter.

(Concentration of Pb: 0.3 ppm)

- Multi-element Standard Solution 0.5 ppm

- Pipette 2.5 mL Multi-element Standard Solution; put into a 50 mL volumetric flask.

- Add 2.5% Nitric Acid solution; to volume and shake until homogenous.

- Filter using a 0.2 µm syringe filter.

(Concentration of Pb: 0.5 ppm)

- Sample Solution

- Weigh 2 g of the sample substance, put it in a 20 mL volumetric flask.

- Add 2.5% nitric acid solution; sonicate until dissolved.

- Add 2.5% Nitric Acid solution to volume and shake until homogenous.

- Filter the filtrate using a 0.2 µm syringe filter.

- Measurement

Follow the Inductively Coupled Plasma Equipment Work Instructions when performing inspections.

  - Assay

    - Requirement :

Molnupiravir Raw Material specification are met a requirement (S-QC-01-280).

    - Procedure :

- Mobile Phase A

- Add 1.2 mL of Ortho-Phosphoric Acid to 1000 mL of Purified Water.

- Stir until homogeneous.

- Filter through a 0.2 µm membrane filter using a vacuum pump.

- Vibrate for 15 minutes in an ultrasonic bath to remove gas.

- Mobile Phase B

- Mix Acetonitril and Purified Water in a ratio of 900 : 100, stir until homogeneous.

- Filter through a 0.2 µm membrane filter using a vacuum pump.

- Vibrate for 5 minutes in an ultrasonic bath to remove gas.

- Solvent

Mix Acetonitril and Purified Water in a ratio of 100 : 900, stir until homogeneous.

- Stock Solution of Impurity D

- Weigh 2.2 mg of standard Impurity D; put it in a 100 mL volumetric flask.

- Add about 50 mL of solvent, sonicate until dissolved.

- Add solvent to volume and shake until homogenous.

(Concentration of Impurity D: 0.022 mg/ml)

- Standard Solution

- Weigh 22 mg of Molnupiravir Working Standard, put it in a 100 mL volumetric flask.

- Add about 50 mL of solvent, sonicate until dissolved.

- Add solvent to volume and shake until homogenous.

- Filter using a 0.2 µm syringe filter; collect the filtrate in the vial then close.

(Concentration of Molnupiravir: 0.22 mg/ml)

- System Suitability Solution

- Pipette 0.5 mL of the Impurity D Stock Solution and 5.0 mL of the Standard Solution into a 50 mL volumetric flask.

- Add solvent to volume, and shake until homogenous.

- Filter using a 0.2 µm syringe filter; collect the filtrate in the vial then close.

(Concentration of Impurity D: 0.22 ppm; Molnupiravir: 22 ppm)

- Sample Solution

- Weigh 22 mg of sample substance, put it in a 100 mL volumetric flask.

- Add about 50 mL of solvent, sonicate until dissolved.

- Add solvent to volume and shake until homogenous.

- Filter using a 0.2 µm syringe filter; collect the filtrate in the vial then close.

- Do duplicate preparations.

- HPLC System

Column	:	Xterra RP18 25 cm x 4,6 mm; 5 m

Column Temperature	:	25C

Sampler Temperature	: 10C

Flow Rate	: 1,0 ml/min

Run Time	:	35 min

Detector	:	UV  260 nm

Injection Volume	:	10 l

Mobile Phase	:	Gradient

| Time (minute) | Mobile Phase A (%) | Mobile Phase B (%) |
| --- | --- | --- |
| 0 | 100 | 0 |
| 5 | 100 | 0 |
| 20 | 60 | 40 |
| 20,1 | 10 | 90 |
| 25 | 10 | 90 |
| 25,1 | 100 | 0 |
| 35 | 100 | 0 |

- Injection Sequence

| No. | Solution Name | Injection Frequency |
| --- | --- | --- |
| 1 | Solvent | 1 |
| 2 | System Suitability Solution | 1 |
| 3 | Standard Solution | 5 |
| 4 | Sample Solution 1 | 1 |
| 5 | Sample Solution 2 | 1 |
| 6 | Standard Solution | 1 |
| 7 | System Suitability Solution | 1 |

- Acceptance Criteria

| Parameter | Acceptance Criteria |
| --- | --- |
| Resolution | Resolution between impurity D - Molnupiravir in System Suitability Solution: not less than 1.5 |
| Area | %RSD area of Molnupiravir peak in Standard Solution: not more than 2.0% |
| Tailing Factor | Tailing Factor of Molnupiravir peak in Standard Solution: not more than 1.2 |

- Calculation

C= LuLs × WsWu × 100%(100% - KA -RS) × B%

C= LuLs × WsWu × 100%(100% - KA -RS) × B%

Note :

Lu	=   Molnupiravir area in the test solution

Ls	=	Average area of Molnupiravir in standard solution

Ws		= 	Weight of the working standard weighed (grams)

Wu	=	Weight of the sample substance weighed (grams)

KA	=	Water Content (%)

RS	=	Amount of Total Residual Solvent (%)

B	=	Assay of Molnupiravir Working Standard as is basis (%)

  - Impurities

    - Requirement :

Molnupiravir Raw Material specification are met a requirement (S-QC-01-280).

    - Procedure :

- Mobile Phase A

- Add 1.2 mL of Ortho-Phosphoric Acid to 1000 mL of Purified Water.

- Stir until homogeneous.

- Filter through a 0.2 µm membrane filter using a vacuum pump.

- Vibrate for 15 minutes in an ultrasonic bath to remove gas.

- Mobile Phase B

- Mix Acetonitril and Purified Water in a ratio of 900 : 100, stir until homogeneous.

- Filter through a 0.2 µm membrane filter using a vacuum pump.

- Vibrate for 5 minutes in an ultrasonic bath to remove gas.

- Solvent

Mix Acetonitril and Purified Water in a ratio of 100 : 900, stir until homogeneous.

- Impurity Standard Solution

- Weigh 22 mg of Molnupiravir Working Standard, put it in a 20 mL volumetric flask.

- Add about 10 mL of solvent, sonicate until dissolved.

- Add solvent to volume and shake until homogenous.

- Pipette 1.0 mL of that solution; put into a 100 mL volumetric flask.

- Add solvent to volume and shake until homogenous (Solution A)

- Pipette 1.5 mL of solution A; put it into a 10 mL volumetric flask.

- Add solvent to volume and shake until homogenous.

- Filter using a 0.2 µm syringe filter; collect the filtrate in the vial then close.

(Concentration of Molnupiravir: 1,65 ppm)

- Stock Solution of Impurity D

- Weigh 2.2 mg of standard Impurity D; put it in a 100 mL volumetric flask.

- Add about 50 mL of solvent, sonicate until dissolved.

- Add solvent to volume and shake until homogenous.

(Concentration of Impurity D: 22 ppm)

- System Suitability Solution

- Pipette 0.5 mL of the Impurity D Stock Solution and 5.0 mL of the Impurity Standard Solution into a 50 mL volumetric flask.

- Add solvent to volume and shake until homogenous.

- Filter using a 0.2 µm syringe filter; collect the filtrate in the vial then close.

(Concentration of Impurity D: 0.22 ppm; Molnupiravir: 0.165 ppm)

- Sensitivity Solution

- Pipette 1.0 mL of solution A; put it into a 50 mL volumetric flask.

- Add solvent to volume and shake until homogenous.

- Filter using a 0.2 µm syringe filter; collect the filtrate in the vial then close.

(Concentration of Molnupiravir: 0.22 ppm)

- Sample Solution

- Weigh 22 mg of Molnupiravir Working Standard, put it in a 20 mL volumetric flask.

- Add about 10 mL of solvent, sonicate until dissolved.

- Add solvent to volume and shake until homogenous.

- Filter using a 0.2 µm syringe filter; collect the filtrate in the vial then close.

- HPLC System

Column	:	Xterra RP18 25 cm x 4,6 mm; 5 m

Column Temperature	:	25C

Sampler Temperature	: 10C

Flow Rate	: 1,0 ml/min

Detector	:	UV  260 nm

Injection Volume	:	10 l

Run Time	:	45 min

Mobile Phase	:	Gradien

| Time (minute) | Mobile Phase A (%) | Mobile Phase B (%) |
| --- | --- | --- |
| 0 | 100 | 0 |
| 5 | 100 | 0 |
| 20 | 60 | 40 |
| 30 | 15 | 85 |
| 35 | 15 | 85 |
| 35,1 | 100 | 0 |
| 45 | 100 | 0 |

- Injection Sequence

| No. | Solution Name | Injection Frequency |
| --- | --- | --- |
| 1 | Solvent | 1 |
| 2 | Sensitivity Solution | 1 |
| 3 | System Suitability Solution | 1 |
| 4 | Solvent | 1 |
| 5 | Standard Solution | 6 |
| 6 | Sample Solution | 1 |
| 7 | Solvent | 1 |
| 8 | Standard Solution | 1 |
| 9 | System Suitability Solution | 1 |

- Acceptance Criteria

| Parameter | Acceptance Criteria |
| --- | --- |
| Resolution | Resolution between impurity D - Molnupiravir: not less than 1.5 |
| s/n | s/n peak in sensitivity solution: not less than 10 |
| Area | %RSD area of peak standard solution: not more than 10.0% |

- The Peak of Impurity

Identify the impurity peak at 0 to 35 minutes, by comparing the chromatogram of the test solution with the solvent (ignore the DMAP peak that appears in the test solution at 5.3 minutes). Calculate the RRT of each impurity that appears and adjust the name of the impurity according to the table below:

| Name | Relative Retention Time (RRT) | Corection Factor (F) |
| --- | --- | --- |
| Impurity A (Ph. Int draft impurity D) | ≈ 0,21 | 0,76 |
| Impurity B (Ph. Int draft impurity A) | ≈ 0,27 | 0,88 |
| Impurity C (Ph. Int draft impurity L) | ≈ 0,85 | 1,15 |
| Impurity D (Ph. Int draft impurity F) | ≈ 1,06 | 0,69 |
| Impurity lain Any other individual impurity | - | 1,00 |
| Molnupiravir | ≈ 1,00 | - |

- Calculation

Calculate the total contamination by adding up all the percentages of impurities that are more than 0.01%.

Cr= AimpAstd × CstdCspl × F × 100%

Cr= AimpAstd × CstdCspl × F × 100%

Note:

Aimp	= 	Area impurity peak in the test solution

Astd	= 	Area impurity peak in the standard solution

Cstd	= 	Concentration of Standard solution (mg/ml)

Cspl	= 	Concentration of test solution (mg/ml)

F	= 	Value of Correction Factor for each impurity

Note:

If after calculating there is an impurity content whose value is less than 0.01%; then reported in the Analysis Report with the editorial "Below LOQ (0.01%)".

  - Residual Solvent - Methode I

    - Requirement :

Molnupiravir Raw Material specification are met a requirement (S-QC-01-280).

    - Procedure :

- Gas Chromatographic System:

Methode	:	GC-Headspace

Column	:	DB-1 60 m x 0,53 mm; 3 µm

Carrier Gas	:	Helium

Flow	:	3,0 ml/menit (constant flow)

Split Ratio	:	2:1

Detector	:	FID

Injection Volume	:	1 ml

Detector Temperature	:	300°C

Injector Temperature	:	250°C

Column Temperature	:	Gradien

| Temperature Increase (°C/min) | Temperature (°C) | Holding Time (min) |
| --- | --- | --- |
| - | 40 | 5 |
| 10 | 230 | 10 |

- Headspace System:

Headspace vial heating temperature	:	80°C

Headspace vial heating temperature-time	:	20 menit

Loop temperature	:	110°C

Transfer line temperature	:	120°C

GC Cycle time	:	43 menit

- Benzene Standard Stock Solution

- Add ± 50 mL of Dimethylsulfoxide into a 100 mL volumetric flask.

- Add 23 µL of Benzene (equivalent to 0.02 g of benzene; ρ = 0.874 g/ml).

- Dilute with Dimethylsulfoxide to volume.

- Shake until homogeneous.

- Standard mixture solution

- Add ± 50 mL of Dimethylsulfoxide into a 100 mL volumetric flask.

- The pipette of each solvent below, put it in the measuring flask above:

| Solvent Name | Volume of Pipette | ρ | Ks |
| --- | --- | --- | --- |
| Methanol | 379 µl (equivalent to 0.3 grams) | 0,792 g/ml | 300,168 mg |
| Isopropyl Alcohol | 639 µl (equivalent to 0.5 grams) | 0,782 g/ml | 499,698 mg |
| Dichloromethane | 45 µl (equivalent to 0.06 grams) | 1,330 g/ml | 59,85 mg |
| Isopropyl Acetate | 575 µl (equivalent to 0.5 grams) | 0,870 g/ml | 500,25 mg |
| Pyridine | 20 µl (equivalent to 0.02 grams) | 0,982 g/ml | 19,64 mg |
| Ethyl Acetate | 554 µl (equivalent to 0.5 grams) | 0,902 g/ml | 499,708 |
| Standard Stock Solution Benzene | 1000 µl | 0,874 g/ml | 874 mg |

* Use different tips for each solvent

- Dilute with Dimethylsulfoxide to volume, shake until homogeneous.

- Standard Solution

- Pipette 10 mL of Standard Mix Solution into a 50 mL volumetric flask.

- Dilute with Dimethylsulfoxide up to volume, shake until homogeneous.

- Pipette 1000 µL of the solution, put it in the Headspace Vial.

- Standard Solution

- Weigh 0.2 g of the sample substance, put it in the Headspace Vial.

- Add 1000 µL of Dimethylsulfoxide, crimp the Headspace Vial.

- Shake until homogeneous.

- Injection Sequence

| No. | Solution Name | Injection Frequency |
| --- | --- | --- |
| 1 | Solvent | 1 |
| 2 | Standard Solution | 6 |
| 3 | Solvent | 1 |
| 4 | Sample Solution | 1 |
| 5 | Standard Solution | 1 |

- Acceptance Criteria for System Suitability Test

| Parameter | Acceptance Criteria |
| --- | --- |
| Resolution | Resolution between Isopropyl Acetate and Benzene peaks > 1.5 |
| %RSD | %RSD peak of standard solutions < 10.0% |

RS= LiLk × Ks100 × 1050 × 1Wu × 100%

RS= LiLk × Ks100 × 1050 × 1Wu × 100%

Calculation:

Note:

RS	=	Residual Solvent (%)

Li	=	Area of each solvent in the sample solution

Lk	=	Average area of each solvent in the standard solution

Ks	=	Equivalence of each solvent (mg)

100	=

1050

1050

Dilution factor of standard solution (ml)

=	Dilution factor of standard solution (ml)

1	=	Dilution factor of sample solution (ml)

Wu	=	Weight of the sample substance weighed (mg)

Notes:

    - If after calculating there is an impurity content whose value is less than the LOQ; then reported in the Analysis Report with the editorial "Below LOQ".

| LOQ Residual Solvent Method I |  |
| --- | --- |
| Isopropanol | 0,005% |
| Diklorometana | 0,0006% |
| Metanol | 0,003% |
| Isopropil Asetat | 0,005% |
| Piridin | 0,001% |
| Etil Asetat | 0,005% |
| Benzena | 0,00002% |

    - If the impurity peak does not appear, then it is reported in the Analysis Report with the editor "Not Detected".

  - Residual Solvent - Methode II

    - Requirement :

Molnupiravir Raw Material specification are met a requirement (S-QC-01-280).

    - Procedure :

- Gas Chromatographic System:

Methode	:	GC-Direct Inject

Column	:	HP-5 30 m x 0,53 mm; 5 µm

Carrier Gas	:	Helium

Flow Rate	:	2,0 ml/menit (constant flow)

Split Ratio	:	2:1

Detector	:	FID

Injection Volume	:	1 µL

Detector Temperature	:	260°C

Injector Temperature	:	150°C

Column Temperature	:	Gradient

| Temperature Increase (°C/min) | Temperature (°C) | Holding Time (min) |
| --- | --- | --- |
| - | 40 | 5 |
| 10 | 150 | 3 |
| 30 | 240 | 8 |

- Standard Stock Solution

- Add ± 50 mL of Dimethylsulfoxide into a 100 mL volumetric flask.

- The pipette of each solvent below, put it in the measuring flask above:

| Solvent Name | Volume of Pipette | ρ | Ks |
| --- | --- | --- | --- |
| Triethylamine | 689 µl (equivalent to 0.5 grams) | 0,726 g/ml | 500,214 mg |
| 2-methyl Tetrahydrofuran | 586 µl (equivalent to 0.5 grams) | 0,854 g/ml | 500,444 mg |
| N,N-dimethylformamide | 93 µl (equivalent to 0.088 grams) | 0,944 g/ml | 87,792 mg |

* Use different tips for each solvent

- Dilute with Dimethylsulfoxide to volume, shake until homogeneous.

- Standard Solution

- Pipette 2 mL of Standard Mix Solution into a 100 mL volumetric flask.

- Dilute with Dimethylsulfoxide up to volume, shake until homogeneous.

- Standard Solution

- Weigh 0.2 g of the sample substance, put into a 10 mL volumetric flask.

- Dilute with Dimethylsulfoxide up to volume, shake until homogeneous.

- Injection Sequence

| No. | Solution Name | Injection Frequency |
| --- | --- | --- |
| 1 | Solvent | 1 |
| 2 | Standard Solution | 6 |
| 3 | Solvent | 1 |
| 4 | Sample Solution | 1 |
| 5 | Standard Solution | 1 |

- Acceptance Criteria for System Suitability Test

| Parameter | Acceptance Criteria |
| --- | --- |
| Resolution | Resolution between Isopropyl Acetate and Benzene peaks > 1.2 |
| %RSD | %RSD peak of standard solutions < 10.0% |

- Calculation

RS= LiLk × Ks100 × 2100 × 10Wu × 100%

RS= LiLk × Ks100 × 2100 × 10Wu × 100%

Note:

RS	=	Residual Solvent (%)

Li	=	Area of each solvent in the sample solution

Lk	=	Average area of each solvent in the standard solution

Ks	=	Equivalence of each solvent (mg)

2100

2100

100	=	Dilution factor of standard solution (ml)

=	Dilution factor of standard solution (ml)

10	=	Dilution factor of sample solution (ml)

Wu	=	Weight of the sample substance weighed (mg)

Note:

If after calculating there is an impurity content whose value is less than the LOQ; then reported in the Analysis Report with the editorial "Below LOQ".

| LOQ Residual Solvent Method II |  |
| --- | --- |
| N,N-dimethylformamide | 0,0044% |
| Triethylamine | 0,05% |
| 2-methyl Tetrahydrofuran | 0,05% |

If the impurity peak does not appear, then it is reported in the Analysis Report with the editor "Not Detected".

  - Impurities 11

    - Requirement :

Molnupiravir Raw Material specification are met a requirement (S-QC-01-280).

    - Procedure :

- Mobile Phase A

- Add 1.2 mL of Ortho-Phosphoric Acid to 1000 mL of Purified Water.

- Stir until homogeneous.

- Filter through a 0.2 µm membrane filter using a vacuum pump.

- Vibrate for 15 minutes in an ultrasonic bath to remove gas.

- Mobile Phase B

- Mix Acetonitril and Purified Water in a ratio of 900 : 100, stir until homogeneous.

- Filter through a 0.2 µm membrane filter using a vacuum pump.

- Vibrate for 5 minutes in an ultrasonic bath to remove gas.

- Solvent

Mix Acetonitril and Purified Water in a ratio of 100 : 900, stir until homogeneous.

- Impurity Standard Solution

- Weigh 22 mg of Molnupiravir Working Standard, put it in a 20 mL volumetric flask.

- Add about 10 mL of solvent, sonicate until dissolved.

- Add solvent to volume and shake until homogenous.

- Pipette 1.0 mL of that solution; put into a 100 mL volumetric flask.

- Add solvent to volume and shake until homogenous (Solution A)

- Pipette 1.5 mL of solution A; put it into a 10 mL volumetric flask.

- Add solvent to volume and shake until homogenous.

- Filter using a 0.2 µm syringe filter; collect the filtrate in the vial then close.

(Concentration of Molnupiravir: 1,65 µg/ml )

- Stock Solution of Impurity D

- Weigh 2.6 mg of standard Impurity D; put it in a 20 mL volumetric flask.

- Add about 10 mL of solvent, sonicate until dissolved.

- Add solvent to volume and shake until homogenous.

(Concentration of Impurity D: 0,13 mg/ml)

- System Suitability Solution

- Weigh 22 mg of Molnupiravir Working Standard, put it in a 20 mL volumetric flask.

- Add about 10 mL of solvent, sonicate until dissolved.

- Pipette 1.0  mL of the Impurity D Stock Solution

- Add solvent to volume and shake until homogenous.

- Filter using a 0.2 µm syringe filter; collect the filtrate in the vial then close.

(Concentration of Impurity D: 6,5 µg/ml; Molnupiravir: 1.1 mg/ml)

- Sensitivity Solution

- Pipette 1.0 mL of solution A; put it into a 50 mL volumetric flask.

- Add solvent to volume and shake until homogenous.

- Filter using a 0.2 µm syringe filter; collect the filtrate in the vial then close.

- Sample Solution

- Weigh 22 mg of Molnupiravir Working Standard, put it in a 20 mL volumetric flask.

- Add about 10 mL of solvent, sonicate until dissolved.

- Add solvent to volume and shake until homogenous.

- Filter using a 0.2 µm syringe filter; collect the filtrate in the vial then close.

(Concentration of Sample Solution: 1.1 mg/ml)

- HPLC System

Column	:	Xterra RP18 25 cm x 4,6 mm; 5 m

Column Temperature	:	25C

Sampler Temperature	: 10C

Flow Rate	: 1,0 ml/menit

Detector	:	UV  230 nm

Injection Volume	:	10 l

Mobile Phase	:	Gradient

| Time (Minute) | Mobile Phase A (%) | Mobile Phase B (%) |
| --- | --- | --- |
| 0 | 100 | 0 |
| 5 | 100 | 0 |
| 20 | 60 | 40 |
| 20.1 | 10 | 90 |
| 25 | 10 | 90 |
| 25,1 | 100 | 0 |
| 35 | 100 | 0 |

- Injection Sequence

| No. | Solution Name | Injection Frequency |
| --- | --- | --- |
| 1 | Solvent | 1 |
| 2 | Sensitivity Solution | 1 |
| 3 | System Suitability Solution | 1 |
| 4 | Solvent | 1 |
| 5 | Standard Solution | 6 |
| 6 | Sample Solution | 1 |
| 7 | Solvent | 1 |
| 8 | Standard Solution | 1 |
| 9 | System Suitability Solution | 1 |

- Acceptance Criteria

| Parameter | Acceptance Criteria |
| --- | --- |
| Resolution | Resolution between impurity D - Molnupiravir: not less than 1.5 |
| s/n | s/n peak in sensitivity solution: not less than 10 |
| Area | %RSD area of peak standard solution: not more than 10.0% |

- The Peak of Impurity

Identify the impurity peak at 0 to 35 minutes, by comparing the chromatogram of the test solution with the solvent (ignore the DMAP peak that appears in the test solution at 5.3 minutes). Calculate the RRT of each impurity that appears and adjust the name of the impurity according to the table below:

| Name | Relative Retention Time (RRT) | Corection Factor (F) |
| --- | --- | --- |
| Impurity 11 | ≈ 0,83 | 1,1 |
| Molnupiravir | ≈ 1,00 | - |

- Calculation

Calculate the total contamination by adding up all the percentages of impurities that are more than 0.01%.

Cr= AimpAstd × CstdCspl × F × 100%

Cr= AimpAstd × CstdCspl × F × 100%

Note:

Aimp			= Area impurity 11 peak in the test solution

Astd			= Area Molnupiravir peak in the standard solution

Cstd			= Concentration of Standard solution (mg/ml)

Cspl			= Concentration of test solution (mg/ml)

RRF	= Relative respon factor of impurity

Note:

If after calculating there is an impurity content whose value is less than 0.01%; then reported in the Analysis Report with the editorial "Below LOQ (0.01%)".

  - Microbial Enumeration Test

    - Requirement :

Molnupiravir Raw Material specification are met a requirement (S-QC-01-280).

    - Procedure :

Follow the Microbial Enumeration Tests Procedure when performing inspections.

- List of Attachment

| Attachments | Document Code | Document Title | Effective Date |
| --- | --- | --- | --- |
| - | - | - | - |

- Revision and Review History

| Document Code | Revision | Effective Date | Changes | Reasons |
| --- | --- | --- | --- | --- |
| I-QC-01-280(E) | 00 |  | - | New Document |

- Distribution List of Document

Original : SSDR

Controlled Copy :

- AMPM

- SPBB